## Sheet1
Important Things to Remember
Plug in ESP32 properly (USB side is facing the side labeled "FRONT"
Do not turn the ON switch on and plug into the ESP32 USB at the same time
Test Name | Test Procedure | Criteria for Success | Completed | Notes
Power rails. | Plug the three main PCBs together. Hook up power supply set to 3.7V to VBATT and turn on main power switch. Measure the power rails. Then sweep power supply input from 3V to 4.2V. | 1V, 2V, and 3V3 rails all are close to their expected value. VIN should be around 3.3V due to reverse diode. +-11V rails should be off. Current consumption looks ok (<100 mA). LOWBATT signal should go low once input is below 3.3V. |  | Current draw is 13 mA at 3.7V. +11V rail is at around 3V, -11v rail is at 0.5V. !V rail is at 0.993V, 2V rail is at 2.004V. LOWBATT triggers at 3.15V.
LiPo | Disconnect power supply, connect LiPo instead. Measure the power rails again. | Voltages are as expected from the previous test.
USB Charging | Connect USB-C to PCB to charge the LiPo. | Voltage rails should turn off. Max charging current should be around 0.8A. Charge LED should turn on (orange). Voltage of battery slowly increases. When charging is done, charge LED turns white.
ESP32 Basic Test | Connect ESP32 and run Blink. | ESP32 blinks the ON LED.
DAC Test | Run ADC_DAC_test on ESP32. Run for every channel of the DAC. | Measured DAC voltages (VGSRAW and VDSRAW) are consistent with what is reported in serial.
ADC Test | Using a breadboard, connect a potentiometer and resistor between the 2V reference rail and GND to create a resistor divider. Feed the output of the resistor divider into VMEAS. Run ADC_DAC_test. | Vary VMEAS by turning the potentiometer around. Probe VMEAS, it should be close to what the ADC is reporting. |  | Bottom PCB has to be disconnected for this test.
DAC and ADC test | Feed the ouput of one of the DAC channels into VMEAS. Run ADC_DAC_test for every channel. | ADC's reported voltage is consistent with what is being written to the DAC. |  | Bottom PCB has to be disconnected for this test. | (positive offset means measured is higher than expected) Channel 0: +20mV, Channel 1: -5mV, Channel 2: +20 mV, Channel 3: no offset, Channel 4: +10 mV, Channel 5: no offset, Channel 6: -5 mV
Gate Voltage Power Rails | Run gate_voltage_test. Test the +-11V power rails. | The +-11V power rails are now on.
Gate Voltage Test, x1 | Run gate_voltage_test with x1 amplification. | Gate voltages measured from a multimeter are consistent with what the gate_voltage_test is reporting in serial.
Gate Voltage Test, x10 | Run gate_voltage_test with x10 amplification. | Gate voltages measured from a multimeter are consistent with what the gate_voltage_test is reporting in serial.
Gate Switch Test | Flip the global gate switch while running gate_voltage_test. | All gate voltages turn off.
Transimpedance Amplifier Test | Solder wires to VDS_BUFF and IDS. Connect these wires to a breadboard with a resistor (>= 1k) in between. Run transimpedance_test, updating test_resistor with whatever resistance was used. | Current expected and current measured are about the same.
Drain Demultiplexer Test | Run multiplexer_test. Test the drain demultiplexer at addresses 0 and 31, measure voltage at VDS1 and VDS32. | With code set to address 0, Vds voltage is observed at VDS1. With code set to address 31, Vds voltage is observed at VDS32.
Source Multiplexer Test | Run multiplexer_test. Test the source multiplexer by choosing addresses 0 and 31 and supplying either IDS1 or IDS31 with an analog voltage (like from a resistor divider). | ADC voltage reported changes as voltage of the chosen mux input changes (not that ADC voltage will not be equal to the mux input voltage)
Demux+Mux Test | Run multiplexer_test.ino. Solder wires to drain 31 and source 31, connect these to a breadboard with a test resistor (>= 1k) in between. (Basically the Transimpedance test but now with the multiplexers in between). | Current expected and current measured are about the same.
Final Test | Run final_boss. Try out Vds sweeps, Vgs sweeps, Vgs amplifcations, etc all at once. | Vgs voltages are what we expect them to be, current readings from transimpedance amplifer are accurate, power rails are stable.
Test Name | Test Procedure | Criteria for Success | Completed | Notes
Boost converter | Hook up boost converter to main PCB assembly. Test heater voltage output. | Before the heater is turned on, the output should be 0V. When heater is turned on, the output is about 7V. Heater ON LED turns on.
Heater | Connect heater to boost converter. Check temperature. (WARNING: Heater gets very hot!) | Heater should rise to about 100C within a few minutes.
Temp Check | Attach thermistor to the side of the heater, and connect the thermistor to the ESP32 Carrier PCB (HEAT THERM header) | ESP32 serial reports correct thermistor readings.
Test Name | Test Procedure | Criteria for Success | Completed | Notes
Humidity Sensor | Connect BME680 to main PCB assembly. Run BME680_test. | Serial reports reasonable temp readings and humidity readings. When you breathe on the sensor, the humidity and temperature should rise.